# THAMWAY PROT series hardware server software
RF Controller section

TCP/IP port to be used
TCP/IP PORT:5027 It uses to communicate with the RF Controller.

About numerical value to be used
Integer
If you put a "0x" at the beginning it is regarded as hexadecimal.
If you put a "0" at the beginning it is regarded as octal.
If you put a "0b" at the beginning it is regarded as binary.
Other than the above it will consider the decimal.
Example： 0x123 → 291
        0b1100 → 12
        0123 → 83
Real
If you put a "u" in the back of the numeric value is regarded as 1e-6(micro).
If you put a "m" in the back of the numeric value is regarded as 1e-3(milli).
If you put a "k" in the back of the numeric value is regarded as 1e+3(kilo).
例: 123k → 123000
   1.2u → 0.0000012
QPSK
At the same time I will output a QPSK pulse and transmitted pulse. There are two 'QPSK1' and 'QPSK2'.
TIME
Unit of time is sec.
FREQUENCY
Unit of frequency is Hertz(Hz).
VOLTAGE
Unit of voltage is volt.

# RFController command list
TCP/IP PORT NUMBER:5027 or callRF()

  1. Common commands
View device information
[Format]
*idn?
[Description]
Make the display of the hardware and software of information connected.
[Result]
THAMWAY,CO477A/B,20100922,DIO2_DISABLE,

Run the Lua script file.
[Format]
run_lua <file name＞
[Description]
Run a lua script file on your PC.
[Example]
run_lua c:/myprograms/test1.lua

Write data to the I/O port
[Format]
outb <I/Oaddress>,<byte data(8bit)>
outw <I/Oaddress>,<word data(16bit)>
[Description]
It is used when operating the register of each device directly.

Read data from the I/O port
[Format]
inb <I/O address>
inw <I/O address>
[Description]
It is used when operating the register of each device directly.
inb: read 8bit data.
inw: read 16bit data.
  1. RF CONTROL COMMAND
Read status
[Format]
STTSR
[Description]

Frequency read and write
[Format]
FREQW<frequency>
FREQR
[Description]
Set frequency MMM.HHHHHH (M=MHz, H=Hz)
[Example]
12.5678MHz FREQW12.567800
98.1MHz FREQW98.100000 (good)
        FREQW98.1 (wrong)
Please specify decimal point always six digits.

TX power level read and write.
[Format]
ATT1W<TX power level>
ATT1R
[Description]
It does the setting of the transmission level.
TX power level 0-1023
[Example]
Set to minimum power	ATT1W0
Set to maximum power	ATT1W1023

Reading and writing of the receiver gain.
[Format]
GAINW<RX gain>
[Description]
Make the settings for the reception gain. The unit is dB.
RX gain 0-95
[Example]
Set to minumum gain GAINW0
Set to maxum gain GAINW95
Reading and writing of the LPFfrequency.
[Format]
LPF1W<LPF frequency>
LPF1R
[Description]
It will read and write LPF frequency. The unit is Hertz.
LPF frequency	100-1000000
[Example]
Set to minimum frequency	LPF1W100
Set to maximum frequency	LPF1W1000000

Reading and writing of the RX phase.
[Format]
PHASW<RX phase>
PHASR
[Description]
It will read and write the receiver phase. The unit is degrees.
RX phase	0.0-359.9
[Example]
Set to minimum phase	LPF1W0.0
Set to maximum phase	LPF1W359.9

Reading and writing of RF SW setting
[Format]
RFSWW<onoff>
RFSWR
[Description]
It will do the setting of TX RF switch.
Onoff	0:RF SW is off.
1:RF SW is on.
[Example]
RFSWW1
RFSWW0